KLA 8300 Vibration Study

# Schneeberger, Inc., 11 DeAngelo Drive, Bedford, MA 01730

8.11.98

## ABSTRACT

KLA has reported excessive vibrations in the Schneeberger XY stage for the 8300 project. The 8300 stage is used for wafer inspection for which precise wafer positioning and repeatability is imperative. The 8300 positioning system consists of the Schneeberger XY stage and a KLA Z stage that is mounted on the top of the XY stage. The inspected wafer sits atop the Z stage platform upon three rubber pads.

The vibration behavior of the positioning system causes the wafer to lose position and actually translate relative to the stage (wafer walk). The vibrations are reportedly caused by the system itself (ballscrew drive); other outside sources of possible vibrations have not been reported.

The purpose of the study is to characterize the behavior of the XY stage and investigate the results of modifying system parameters in an attempt to improve positioning performance.

### INTRODUCTION

There are several ways to control unwanted vibration behavior in a system. Depending on the system’s configuration and constraints, certain vibration control means are preferable.

- Because the forcing frequency is generated by the drive, and the drive is an integral part of the system, the vibration source can not be isolated. If there are other separate sources generating forcing frequencies in the system (pumps, etc.) isolation of the source should be investigated.

- The amplitude of the vibrations generated by the ballscrew may be adjusted by controlling the idling torque and preload.

- By increasing the forcing frequency sufficiently, system damping may be accomplished. Driving the ballscrew at faster speeds will generate forces at higher frequencies and will have a reduced excitation effect on the system.  Increased velocities are likely to be accompanied by increased accelerations and inertial forces. These inertial forces might lead to wafer slip.

- By increasing the mass of the system, damping is accomplished. The natural frequency is inversely proportional to M½ . A higher mass results in a lower natural frequency that is more resistant to excitation. Drawbacks of increased mass include cost considerations and motor drive system implications because of the increased inertia.

- By decreasing the stiffness of the system, damping is accomplished. The natural frequency is directly proportional to (Stiffness)½. A lower stiffness results in a lower natural frequency that is more resistant to excitation. But since precise movement depends upon stiff rolling elements, a lower stiffness is in direct conflict with the stages important performance requirement of high stiffness. In addition, damping elements placed in intermediate locations of the system (under the Z stage, where it mounts to XY stage) are not feasible because the accompanying hysterisis results in inaccurate positioning.

- The way in which the system reacts to positioning error is controlled by the control parameters. The optimum control parameters are based on the forcing function and the system’s physical parameters. Because of the complexity of the system, the control parameters can not be found analytically through a mathematical model; instead the system must be tested and fine-tuned. Because the system’s parameters can actually change (location of masses during operation, ballscrew idling torques, etc.) a set of control parameters for optimum performance must be found to encompass the complete band of such changes.

Badly tuned control parameters may induce new forcing frequencies into the system.  The way the system responds to positioning error is adjusted by these parameters. An oscillatory response could result if the system continually overcompensates for the positioning error. The ballscrew would essentially be vibrating the table because of the controller’s commands.

- See #1

The amount of force that is transmitted from a vibration source to a system is dependent on the ratio between the forcing frequency and the systems natural frequency. The best case is a very high ratio; a very high forcing frequency and/or a very low natural frequency. Increased system mass and/or decreased system stiffness result in a lower natural frequency, thus the frequency ratio increases and transmissivity decreases. The worst case is a frequency ratio of 1. Under such conditions resonance occurs resulting in very large and damaging oscillations.

If the system’s parameters and the forcing frequency are near resonant conditions, means to increase the frequency ratio should be investigated i.e. increase mass, decrease stiffness, increase forcing frequency. This would result in the force being dampened. If this is not possible, means to decrease the ratio should be investigated. Although this will not provide dampening (full transmission of force is the best performance that this method can achieve) it will reduce the damaging oscillations and forces which accompany resonance.

Test  Setup

The XY stage was bolted down to a mounting table. Servo motors where connected to drive the ballscrews in the X and Y directions. X and Y axes where equipped with linear encoders for feedback. Encoder feedback and motor controls where wired to computer controls. PID control was implemented for table positioning. The Z stage was mounted on the XY stage. X, Y and Z accelerometers where attached to the Z stage to acquire data. A spectrum analyzer processed the data to correlate power and frequency for each axis. A magnifying scope was set up above the wafer. The scope was connected to a monitor to view the wafer. A grid generator was used to pinpoint a specific location on the wafer.

Test Procedure

Three XY stages were subjected to the following tests. The first tests were done with a set of standard parameters provided by KLA:

Velocity = 80 mm/sec

Acceleration = 320 mm/s2

Jerk = 1280 mm/s3

##### Velocity profile = S-curve

Control Parameters:

##### X:	Kp = 65		Ki = 3		Kd = 140

Y:	Kp = 80		Ki = 3		Kd = 280

Variations in performance where studied by varying the control parameters (Kp, Ki, Kd), velocity, velocity profile, acceleration, and jerk.

Testing was done with amplifier and controls provided by KLA for the X and Y axis.

Testing was also done with Etel motor, amplifier, and controls for the X axis.

#### Wafer Walk Test

The wafer was positioned on the Z stage atop the XY stage. A unique feature of the wafer was found and pinpointed using the cross hairs. The XY stage was then driven with the following positioning sequence:

X (mm)	Y (mm)

0	0

-120	180

0	0

120	180

0	0

-120	-180

0	0

120	-180

0	0

-120	0

0	0

120	180

0	0

At the end of the positioning sequence, wafer position was checked.

This test was done with the wafer on rubber pads and on metal washers.

#### Vibration Test

Three accelerometers (X, Y, and Z) were positioned on the Z stage. The three accelerometers where then placed on the XY stage without the Z stage attached. The following frequency and vibration power data were measured:

- Resonant frequency

- Power at resonant frequency

- Energy (“bandpower”) in frequency band 0 – 400 Hz

The above quantities were measured during the following positioning strokes:

X = 120mm

X = -120mm

##### Y = 120 mm

Y = -120 mm

Natural Frequency Test

The following setups were struck with a small hammer to impart an impulse and generate a frequency response:

-Z stage on XY stage

-Z stage alone

-XY stage alone

Using the spectrum analyzer, bandpower was plotted versus frequency.

###### Experimental Results

Wafer Walk

Rubber Pads –

The wafer was tested for wafer walk under several different parameters. The wafer mounted on the rubber pads did not walk under the following conditions:

- Standard KLA parameters

- Increased velocity

- Increased accelerations

- Some resonance.

- Setup with Etel amplifier and controls.

The wafer slipped only under high resonance.

Metal Washers-

The wafer mounted on the metal washers experienced a drift of X = .6m and Y = 1.3 m under the following conditions: The standard KLA parameters run through the “wafer walk test” cycle five times

The wafer was more susceptible to walking than when mounted on the rubber pads. High resonance and large accelerations, especially with a trapezoidal ramp velocity profile, caused large wafer slippage. Under tuned control parameters, wafer positioning was repeatable.

Vibration Test

Vibration tests were done under several different parameters. In addition to using KLA’s standard parameters, the proportional and derivative gains were adjusted. Under resonance, the stage exhibited high bandpower and high peaks. By adjusting the gains, the resonance was prevented and lower bandpower and peaks where measured. Under higher speeds, accelerations, and jerk, bandpower also decreased and the peaks flattened out substantially.

With the Etel setup on the X-axis, the control parameters were adjusted to test performance. Under a large range of parameters, resonance could not be generated in the stage. Both bandpower and peak power levels were reduced as compared to the KLA drive system.

Discussion

The behavior of the XY stage was greatly influenced by changes in ballscrew speed, proportional gain and derivative gain. At higher speeds, the ballscrew generated forces at higher frequencies. These higher frequencies are more difficult to transmit to the system and therefore the stage is less apt to resonate. By adjusting the proportional and derivative gains, resonance could also be avoided. Badly tuned control parameters could make the stage resonate loudly. An adjustment of the proportional and derivative gains eliminated the resonance.

The decreased vibration behavior was indicated by decreased bandpower and lower peak power heights. The bandpower represents the total power of the stage’s vibrations summed between 0 Hz and 400 Hz. At certain frequencies, peaks occurred indicating relatively high vibration behavior. It is theorized that wafer walk is most likely to occur at such frequencies where large amounts of power are expended in stage vibration. Large vibrations in the Z direction could results in the wafer losing contact with the support pads and therefore losing position. Vibrations in the X and Y directions could provide tangential forces on the support pads high enough to overcome the static friction and cause wafer walk.

With the Etel setup, the stage did not resonate for a wide range of PID parameters. It is possible that the smaller available torque output of the Etel motor as compared to the motors used by KLA prevented overcompensation for error control.

The very good wafer positioning repeatability even under high vibrations did not match the stage performance as reported by KLA. This is the result one of the following:

- The three stages tested are better stages.

- The testing setup does not fully reflect the operating environment of KLA.

- KLA did not implement the optimum control parameters for the system.

For #1, there is a significant variation in the ballscrew idling torques between the three tested stages and the earlier prototype stages. The torque variation for the tested stages (2 – 3.5 N-cm) is lower than that for the prototype stages (3 – 6 N-cm).  This could have an effect on the amplitude and frequency of the forces generated by the ballscrew as controlled by the PID controls.

In reference to #2, the most notable difference between the two setups is that KLA operates the stage in a vacuum chamber. The Schneeberger test setup was operated under standard atmospheric conditions. Thus any effects of a vacuum on the system were not reflected in the test setup and corresponding performance results. Changes such as variations in the rubber pad properties (coefficient of static friction) would result in different wafer slip performance. To test the effect of a lower coefficient of static friction between the wafer and the mounts, the rubber pads were replaced with metal washers.  As expected, there was an increased susceptibility to wafer walk. Nevertheless, position repeatability was still achieved with the metal washers as wafer mounts when the PID controls were properly set.

Finally, it is possible that the stages in use by KLA are not operating with the best control parameters. If the control parameters are fine tuned, improved stage performance could result.

Conclusion

The ultimate performance criterion, wafer positioning repeatability, was achieved with all three tested XY stages. Even with high vibrations, the wafer did not slip. Only under extreme resonance did the wafer slip.

With properly set control parameters, the XY stage performed very well. Excessive vibrations could be avoided and wafer slippage did not occur, even under high accelerations. Based on future experimental tuning, one set of control parameters should be determined that will encompass slight variances in the stage behavior. Such variations include ball screw idling torque and center of mass location changes during the stage’s movement.

Decreased vibrations also resulted from increased ballscrew speeds. With higher driving speeds, the forces are generated at higher frequencies. These higher forcing frequencies are more difficult to transmit to the system.

If KLA does not confirm the performance results of the three XY stages, any other differences between the test setup and the actual operating environment should be investigated to pinpoint all possible causes for the reported decreased performance.

Finally, if the desired stage performance is still not achieved, Schneeberger would recommend taking over full responsibility for the complete drive system (motor, amplifier, controls, cables). With higher-grade electronics, the desired performance is more readily achievable.

# Text boxes

Desired Position

Figure 1

Actual Position